## Slide 1: Non-invasive and pulse-resolved polarization diagnostics
Joakim Laksman
X-ray Photon Diagnostics (XPD)

## Slide 2
Acknowledgements
European XFEL: XPD
  Jan Grünert
  Marc Planas
  Florian Dietrich
  Jia Liu
  Theophilos Maltezopoulos
  Wolfgang Freund
  Torben Falk
  Sonay Sayar
European XFEL: SQS
  Marcus Ilchen
DESY: PETRA III: P04
  Jens Viefhaus
  Jens Buck
  Leif Glaser
  Frank Scholz
  Gregor Hartmann
European XFEL science support groups : VAC, CAS, AE, ITDM, EET
DESY : Transport, Survey&Alignment
Laksman et al., J. Synchrotron Rad., Accepted Manuscript, (2019)

## Slide 3: Photo-Electron Spectrometer Overview
16 eTOF spectrometers perpendicular to X-ray beam.
  TOF spectra  Photon energy
  Angular distribution  Polarization
Gas targets
  N2
  Ne
  Kr
  Xe
X-ray pulse ionizes gas in extraction region.
Photo electrons hit the 16 MCP based detectors which gives time-of-flight spectra on digitizers
Helmholtz-coil & Magnetometer
Movement of entire system with six degrees of freedom.

## Slide 4
Measurements at SASE3

## Slide 5
Spectra provides two types of information
TOF  Kinetic energy  Photon energy
Angular distribution  Polarization

## Slide 6
Polarization
PLin = 0.96 ± 0.03
y = 0.02 ± 0.03 [rad]
For every pulse we collect 16 TOF spectrums in all directions.
Spectra provides two types of information
TOF  Kinetic energy  Photon energy
Angular distribution  Polarization
Fitting procedure provides us with polarization.
Future afterburner will produce circular polarization
Polarization important for many experiments at end-station

## Slide 7
Polarization
Unreliable pulse triggers Veto-signal
Instrument does not store data for that pulse
Saves storage space on server
Compare to multilayer based polarimeter
Advantages
  Pulse resolved
  Non-invasive
Dis-advantages
  Only linear information
  Cannot distinguish between un-polarized or right / left polarization

## Slide 8
N2 KVV Auger line for normalizing detector efficiency.

## Slide 9
J. J. YEH and I. LINDAU, AT. DATA NUCL. DATA 32 (1985)
Neon
Orbital BE [eV]
Select suitable
target gas:
Cross Section
Criteria
High cross section (but still avoid space charge)
Anisotroy parameter b far from 0
Kinetic energy as low as possible

## Slide 10
Status:
In principle operational:
  All hardware is installed in tunnel.
  Software enables us to acquire and analyse data.
  Spectrometer has been calibrated: TOF can be converted to Photon-energy. Angular distribution can be converted to Polarization.
Ongoing work and outlook:
Only collect raw data trace which is storage consuming and requires post-analysis.
PES data source cannot be included in experimental end-stations.
In collaboration with CAS and AE we are developing MDL-Device to convert huge data trace to relevant parameters in real time:
  Spectral distribution
  Polarization
Purpose of data:
  Data must be stored in .h5 files together with user experiments.
  Continuously available to operators for tuning undulator gap and variable polarization.

## Slide 11
Status:
In principle operational:
  All hardware is installed in tunnel.
  Software enables us to acquire and analyse data.
  Spectrometer has been calibrated: TOF can be converted to Photon-energy. Angular distribution can be converted to Polarization.
Thank you very much for your attention!
Ongoing work and outlook:
Only collect raw data trace which is storage consuming and requires post-analysis.
PES data source cannot be included in experimental end-stations.
In collaboration with CAS and AE we are developing MDL-Device to convert huge data trace to relevant parameters in real time:
  Spectral distribution
  Polarization
Purpose of data:
  Data must be stored in .h5 files together with user experiments.
  Continuously available to operators for tuning undulator gap and variable polarization.